# Analytical Method Validation Protocol

Prepared in accordance with ICH Q2(R2): Validation of Analytical Procedures

| Protocol No. | ________________ |
| --- | --- |
| Product / Method Name | ________________ |
| Analytical Technique | ________________ (e.g., HPLC, GC, UV-Vis) |
| Prepared By / Date | ________________ |
| Reviewed By / Date | ________________ |
| Approved By / Date | ________________ |

# 1. Purpose and Scope

This protocol defines the experimental design, acceptance criteria, and documentation requirements for validating the above analytical procedure, in accordance with ICH Q2(R2) "Validation of Analytical Procedures: Text and Methodology." Validation demonstrates that the method is suitable for its intended purpose prior to routine use.

This protocol applies to the validation characteristic(s) applicable to the method category: identification, assay (content/potency), or impurity testing (quantitation and limit).

# 2. Reference Documents

ICH Q2(R2): Validation of Analytical Procedures

USP General Chapter <1225>: Validation of Compendial Procedures

Internal SOP for [instrument/technique]

# 3. Validation Characteristics

## 3.1 Specificity

Objective: Demonstrate the method's ability to unequivocally assess the analyte in the presence of components that may be expected to be present (impurities, degradants, matrix, placebo).

Acceptance Criteria: No interference from blank/placebo/degradation products co-eluting or overlapping with the analyte peak; peak purity acceptable.

Procedure: [Describe experimental design, number of replicates/levels, and sample preparation specific to this method.]

| Trial / Level | Result | Acceptance Criterion | Pass / Fail |
| --- | --- | --- | --- |

Conclusion: [ ] Meets acceptance criteria    [ ] Does not meet acceptance criteria — see deviation report

## 3.2 Linearity

Objective: Demonstrate that test results are directly proportional to analyte concentration across the validated range, typically evaluated at a minimum of 5 concentration levels.

Acceptance Criteria: Correlation coefficient (r) ≥ 0.999 (or per method-specific requirement); y-intercept and residuals evaluated for systematic deviation.

Procedure: [Describe experimental design, number of replicates/levels, and sample preparation specific to this method.]

| Trial / Level | Result | Acceptance Criterion | Pass / Fail |
| --- | --- | --- | --- |

Conclusion: [ ] Meets acceptance criteria    [ ] Does not meet acceptance criteria — see deviation report

## 3.3 Accuracy

Objective: Demonstrate closeness of agreement between the accepted reference value and the value found, typically evaluated at a minimum of 3 concentration levels with 3 replicates each (9 determinations minimum).

Acceptance Criteria: % Recovery within 98.0%–102.0% (or per method-specific requirement).

Procedure: [Describe experimental design, number of replicates/levels, and sample preparation specific to this method.]

| Trial / Level | Result | Acceptance Criterion | Pass / Fail |
| --- | --- | --- | --- |

Conclusion: [ ] Meets acceptance criteria    [ ] Does not meet acceptance criteria — see deviation report

## 3.4 Precision — Repeatability

Objective: Evaluate agreement between replicate measurements under the same operating conditions over a short interval of time, minimum 6 determinations at 100% of test concentration, or 9 determinations across the specified range.

Acceptance Criteria: %RSD ≤ 2.0% (or per method-specific requirement).

Procedure: [Describe experimental design, number of replicates/levels, and sample preparation specific to this method.]

| Trial / Level | Result | Acceptance Criterion | Pass / Fail |
| --- | --- | --- | --- |

Conclusion: [ ] Meets acceptance criteria    [ ] Does not meet acceptance criteria — see deviation report

## 3.5 Precision — Intermediate Precision

Objective: Evaluate the effect of within-laboratory variations (different days, analysts, or equipment) on precision.

Acceptance Criteria: %RSD ≤ 2.0% between analysts/days (or per method-specific requirement).

Procedure: [Describe experimental design, number of replicates/levels, and sample preparation specific to this method.]

| Trial / Level | Result | Acceptance Criterion | Pass / Fail |
| --- | --- | --- | --- |

Conclusion: [ ] Meets acceptance criteria    [ ] Does not meet acceptance criteria — see deviation report

## 3.6 Detection Limit (LOD)

Objective: Determine the lowest amount of analyte that can be detected, but not necessarily quantitated, using the signal-to-noise or standard-deviation-of-response/slope approach (3.3σ/S).

Acceptance Criteria: LOD calculated and reported with supporting raw data; signal-to-noise ≥ 3:1 at the reported LOD concentration.

Procedure: [Describe experimental design, number of replicates/levels, and sample preparation specific to this method.]

| Trial / Level | Result | Acceptance Criterion | Pass / Fail |
| --- | --- | --- | --- |

Conclusion: [ ] Meets acceptance criteria    [ ] Does not meet acceptance criteria — see deviation report

## 3.7 Quantitation Limit (LOQ)

Objective: Determine the lowest amount of analyte that can be quantitatively determined with suitable precision and accuracy, using the standard-deviation-of-response/slope approach (10σ/S).

Acceptance Criteria: LOQ calculated and reported; signal-to-noise ≥ 10:1; precision and accuracy confirmed at the LOQ concentration.

Procedure: [Describe experimental design, number of replicates/levels, and sample preparation specific to this method.]

| Trial / Level | Result | Acceptance Criterion | Pass / Fail |
| --- | --- | --- | --- |

Conclusion: [ ] Meets acceptance criteria    [ ] Does not meet acceptance criteria — see deviation report

## 3.8 Range

Objective: Establish the interval between the upper and lower concentrations of analyte for which acceptable precision, accuracy, and linearity have been demonstrated.

Acceptance Criteria: Range confirmed to span at minimum 80%–120% of the target/specification concentration (assay) or LOQ to 120% of specification (impurities), per ICH Q2(R2) Table 1.

Procedure: [Describe experimental design, number of replicates/levels, and sample preparation specific to this method.]

| Trial / Level | Result | Acceptance Criterion | Pass / Fail |
| --- | --- | --- | --- |

Conclusion: [ ] Meets acceptance criteria    [ ] Does not meet acceptance criteria — see deviation report

## 3.9 Robustness

Objective: Evaluate the method's capacity to remain unaffected by small, deliberate variations in method parameters (e.g., mobile phase composition, flow rate, column temperature, pH).

Acceptance Criteria: System suitability and result acceptance criteria still met under each varied condition tested.

Procedure: [Describe experimental design, number of replicates/levels, and sample preparation specific to this method.]

| Trial / Level | Result | Acceptance Criterion | Pass / Fail |
| --- | --- | --- | --- |

Conclusion: [ ] Meets acceptance criteria    [ ] Does not meet acceptance criteria — see deviation report

# 4. System Suitability

System suitability testing shall be performed prior to and/or during sample analysis per USP <621>, confirming resolution, tailing factor, theoretical plates, and %RSD of replicate injections meet the method-specific criteria (see the companion HPLC System Suitability Checklist template).

# 5. Deviations and Out-of-Specification Results

Any result failing to meet the acceptance criteria defined above shall be documented via a deviation/OOS investigation per site quality procedures before the method is declared validated.

# 6. Validation Summary and Approval

Overall Conclusion: [ ] Method is validated for its intended use    [ ] Method requires revalidation/further work

| Role | Name | Signature | Date |
| --- | --- | --- | --- |